Towards high  confidence medical device  software

Zhihao Jiang,  Houssam Abbas, Kuk  Jin Jang and Rahul  Mangharam

Department of Electrical and Systems Engineering, University of Pennsylvania, USA

1   The  Promise and Challenges of Medical Devices

Medical  devices  play  an essential role in the care of patients  around  the world, and can have a life-saving effect. To cite one example, an estimated 3 million peo- ple worldwide  have implanted  pacemakers  (a heart rhythm  adjustment  device), with ˜600,000 added annually.  In the US, 8˜00,000 people have an implanted defib- rillator (another heart rhythm management device),  with 10,000 added monthly. Clinical  trials have presented evidence that patients implanted with defibrillators have a mortality  rate reduced  by up to 31%.

Financially, the medical device market is worth $289 billion. Of that, $110 bil- lion is in the US alone, with this number pro jected to reach $133 billion in 2016. Examples include everything from adhesive bandages to drug infusion pumps, surgical robots, deep brain stimulation systems and devices still undergoing basic research like the artificial pancreas. These are safety-critical technologies com- bining hardware  and software, each of which must be rigorously verified to be efficacious and safe.

According to  the  US  Food  and  Drug  Administration,  in  1996, 10%  of all medical  device  recalls  were  caused  by  software-related  issues.  This  percentage rose to an average  of 15% of recalls from 2008 to 2012. Implanted  cardiac  pace- makers  and  defibrillators  have  approximately  80,000-100,000 lines  of software code [3] which  essentially  makes all sensing, control and actuation  decisions autonomously  within  the  human  body,  over  the  5-7 year  device  lifetime.  The primary  challenge of high-confidence medical  device software is to guarantee  the device  will  never drive  the patient  into an unsafe condition even though  we do not have complete  understanding  of the physiological  plant.

1.1    Human-in-the-loop Medical Devices

We  may  classify  devices  as  open-loop  or closed-loop.  A  closed-loop device like

a  pacemaker  is  in  a  feedback  loop  with  the  organ(s)  it  effects (see  Fig.1):   it monitors  certain  physiological   variables   like  heart  rate,  and  delivers  therapy,

in  the  form of low-energy  electrical  pulses,  to  maintain  a  healthy  heart  rate. Another  example  is the artificial  pancreas,  which  monitors blood glucose  levels and delivers  therapy,  in the form of insulin, to maintain  safe glucose  levels.  An

open loop device on the other hand, either, (a) like a drug infusion pump, does not

measure  any  physiological  variables:  the therapy  it delivers  is pre-programmed and non-reactive;  or (b)  only measures physiological  signals but does not deliver

therapy,  as in the case of a blood pressure monitor.

Closed-loop  devices  require  very  little  physician  intervention  after  the  dis- charge visit, and hence permit a better lifestyle. Because they are constantly monitoring  the  physiological   variables,   they  permit  a  more  timely  delivery  of therapy.  The  complex  run-time  diagnoses  needed for closed  loop performance, and the intricate therapy  delivered,  has driven most diagnosis and therapy  func- tions into software. This  software is life-critical  and verification  methods should provide  a high confidence in its correctness.

Pacemaker Pulse

Ventricular

lead signal

Fig. 1. Pacemaker operating in a closed-loop with the  heart. The leads sense  cardiac electrophysiological activity from  inside the  heart tissue (AS/ VS  = Atrial/Ventricular Sense event) and actuate the  heart (AP/ VP  = Atrial/ Ventricular Pacing event

to maintain a desired heart rate.

Verifying the safety and  efficacy of  closed-loop software,  by  definition,  re­ quires that the device be connected to  the organ(s) it is affecting. For example, in  the case of a pacemaker, that would be  the heart of a living patient. But  with the advent of computer models of physiological functions, such as  those encom­ passed by  the Physiome project or presented later in this article, the model-based design  ( MBD) of closed-loop  medical  devices  presents  efficient  complementary approaches that are actively researched in various disciplines of  engineering and computer science. In  MBD, the device (or  a model thereof)  is  connected to  a model  of  the "physiological plant"  it  interacts with. By  high  confidence verifi­  cation, researchers mean that  under all  possible behaviors  of  the  physiological models, the device will  act  correctly.

There are two  major  differences between modeling physiology and  model­ ing   man-made systems: first, physiology is  much more complex and  less  well­  understood than man-made systems like  cars and  airplanes, and  spans several scales from the molecular to the entire human body. Secondly, the variability be­  tween humans is orders of magnitude larger than that between two  cars coming off the assembly line.  Using the pacemaker as  an  example  of  closed-loop device, and the heart as  the organ to  be  modeled, we  present  several of the challenges and early results in  model-based verification.

2   Life-critical Closed-loop Software

The  heart  is a  specialized  muscle  that  pumps  oxygenated  blood  to the rest of the body.  It is composed  of four chambers:  two upper  chambers  called  the left and right atrium,  and two lower chambers called  the left and right ventricle,  are synchronized respectively.  In a healthy  resting adult,  the heart rate is 60 to 100 beats  per minute  (each  ventricular  contraction  is a  beat).  The  contractions  of the heart are controlled by the waves of spontaneous electric depolarization  that

traverse it regularly.  A spontaneous electric current originates in the Sino-Atrial

(SA)  node in  the  right  atrium  and  propagates  throughout  the  atria,  causing

them to contract. It then propagates down to the ventricles along well-defined conduction  pathways,  causing the ventricles  to contract  in turn. The  SA  node is thus termed the natural  pacemaker  of the heart.

Under certain diseased conditions, the heart rate drops below what is needed to maintain  adequate  blood  flow to the  body.  This  clinical  condition  is called

bradycardia. When  such  a  heart  rate  drop  is due  to abnormalities  in the elec-

trical conduction system, an implanted pacemaker might be recommended as treatment.  A pacemaker  is implanted  near the left collar of the patient as shown in Fig.1,  and has two leads:  one connects  to the right  atrium,  the other to the

right ventricle.

The leads act as both sensors and effectors: if the pacemaker  senses a missing electric  spike on either lead,  indicative  of a missed contraction,  it will  send an electric signal to the corresponding chamber  to provoke contraction,  thus acting as an artificial pacemaker.  The algorithms for detecting missed beats are complex and implemented in software which runs on the pacemaker  itself. Part  of the difficulty of performing that detection comes from the great variability in heart rates between patients and indeed within a single patient across time. Moreover, because  the pacemaker  is limited  to sensing electrical  activity through  its two leads, different phenomena can manifest themselves identically  to the pacemaker, thus making  detection  even harder.

In  Endless  Loop  Tachycardia (ELT),  this  ambiguity   causes  the  pacemaker

to actually  induce dangerously  elevated  heart  rates (tachycardia), which  would

not have  arisen  had  the  heart  been  operating  on its  own.  This  is an  example

of a adverse closed-loop condition : a dangerous  situation  that  arises as a result

of the  interaction between  heart  and  device.  No  amount  of open-loop  device

testing and verification can reveal this condition - hence the need for closed-loop

validation  of medical  devices,  and  for physiological   heart  models  that  enable early  and affordable closed-loop validation.

3   Choosing the right model  for the job

Heart models of different kinds have been developed  for a range of applications and Fig. 2 shows four heart modeling approaches that emphasize the electric, mechanical,  cellular  and fluid flow aspects  of cardiac  function.  Several  of these modeling  approaches  employ  over 4 million finite elements or 100,000 ordinary differential equations to describe the dynamics and take several hours to simulate a single cardiac  cycle.

Cellular models describe  the generation  and  spread  of electrical  action  po-

tentials (i.e., voltages)  at the moleculo-cellular level [5]. At  the cellular  level, the flow of charged  ions into and out of the cardiac  cell is responsible for the change

in voltage  across  the  cell  membrane.  Cellular   models  of electrical  activity  are used to study  how activity across ion channels affect the relation  between elec- trical and mechanical  behaviors of heart tissue, as well as to study drug therapies that affect the ion channels properties.

Anatomical models are developed  using imaging  technologies  like MRI,  and

seek to re-create detailed anatomical  structures like fiber orientations and the distribution  and extent  of scar  tissue. These  structures  affect the heart’s  oper- ation by determining muscle contraction and modifying the speed and paths of electrical conduction throughout the heart. Thus anatomical models provide a foundation  for whole heart  modeling  efforts that  we cover  next.  They  are  also used to simulate  the effects of certain  medical  devices  like stents and  artificial valves.

Whole heart models use a  continuum  approximation  of the cellular  models

of electrical  propagation  with the structure obtained from anatomical  models. Researchers  have  developed  Partial Differential  Equations  models  of electrical

activity in the whole heart to analyze  the mechanisms  of various  arrhythmias.

Researchers  at Johns Hopkins [6] further used these models to predict  the onset

of arrhythmias  and propose potential  therapies.  Such electro-mechanical models

help evaluate  the mechanical  effects of different arrhythmias  on blood flow.

Electro-Physiological (EP) heart  models  help  study  the  timing  properties

of the generation and propagation  of electrical  signals through the electrical conduction system, and can accurately diagnose most arrhythmia.  Unlike the previous  heart models which  are built  from the cellular  level on up, EP  models

are developed  by  conducting  a clinical  EP  testing procedure,  and are amenable

to model checking [1], a powerful  verification  technique  pioneered  in the semi-

conductor  industry.  EP  testing  is a common method  to diagnose  arrhythmias: a  physician  inserts  catheters  with  electrodes  into  the  patient’s  heart  through the veins and  measures  the local  electrical  activity around  the electrodes.  The

physician  uses the patterns of electrical  activity and its timing characteristics  to

diagnose the heart’s  condition.  In particular, electrical  timing parameters  of ac- tion potentials in a tissue region like conduction delay, rest period and refractory period are measured,  and any abnormal  conduction  paths are detected.

In  [2], researchers  from  the  University   of Pennsylvania  developed  a  heart

model based on clinical  EP  testing. Since the pacemaker  only looks at the timing

of events  as  input,  this  EP   model  only  seeks  to  model  the  correct  timing  of

Cellular Model

Anatomical Model

Electrophysiology Model

Electro-mechanical Model

Fig. 2. Modelling different phenomena in the heart

electrical  activity in select tissue of the heart. Specialized  tissue like the SA  node generates electric  events and is modeled by  a finite timed automaton  known as

node automaton. The  node automaton  models the timing  of signal  generation,

blocking  and transmission in heart structures  like the AV  node. The  rest of the tissue is abstracted  as variable  conduction delays as path automata between node automata.  Different heart conditions can be modeled by varying  the number and connectivity  of note/path  automata  and their parameters (Fig.  3.(b)).  Moreover, pacing  applied  to the heart  can be represented  as an external  activation  signal to the node automata.  EP  models have been validated  by  physicians,  and have been used for model checking  pacemaker  software.

AS	AS

SA node

AV node

VS 	VS

His-Purkinje System

(a) 	(b) 	(c)

Fig. 3. Timed automata (b) capture the timing relations between electrical events in the heart, and can produce sequences of sense events (c) to the pacemaker.

4   Closed-loop model  checking  of device  software

Short  of clinical  trials  (which  we cover  later  in this article),  current  validation practice  for closed-loop medical devices focuses on open-loop testing and reviews of the design  process.  In testing,  a set of input  sequences  is fed to the device, and the device’s  output  is checked  for correctness,  typically by  comparing  it to a pre-defined expected  output.  Such open-loop testing does not test the effect of the device  on the organ: e.g., we can’t  test how the heart rate changes following a pacing  by the pacemaker.  Thus,  we need a heart model that can interact  with the device. If we use, say, a high-fidelity PDE-based model such as the electro- mechanical models described above to react to the device and generate input sequences, there are an infinite number of heart rhythms  that such a model can generate,  and testing only uses a finite subset of those. Thus,  testing (whether open-loop or closed-loop with a heart model) is necessarily an incomplete tech- nique, and safety violations  of the device  may  be missed during testing.

The  timed  automata-based  EP   models  are  amenable  to  model checking, a

technique  that  mathematically  explores  all  possible  executions  of  the  heart

model and device software combination against specified requirements (e.g. the pacemaker will not pace the heart beyond an upper rate limit). Model checking is widely used in the semiconductor industry to verify chip designs at various levels of abstraction,  in particular  at the Register  Transfer Level  (RTL). Violations  of the requirements are returned by the model checking tool as an execution trace, which can be analyzed  and used to improve the system. To capture  the variabil- ity  in the heart’s  behavior  (more  generally,  in the physiological  phenomena  of

interest),  the heart  model is non-deterministic: for example,  rather  than speci-

fying that the conduction  delay  in the AV  node is always  0.14ms, we allow it to be any value in the correct physiological  range [0.12, 0.2] ms. The  model checker will symbolically explore all executions corresponding to all values in this range

(rather  than  select  a  few)  in search  for requirements  violations.  Subtle  errors in the design of safety-critical systems that  often elude conventional  simulation and testing techniques can be (and have been) found in this way.  Because  it has been proven cost-effective and integrates well with conventional design methods, model checking  has being  adopted  as a standard  procedure  for the quality  as- surance  of automotive  and  avionics  systems,  but  has yet  to enter the world  of medical  devices.

Endless  Loop  Tachycardia (ELT)  is  one  example  of  a  safety  hazard  that arises  in  the  interaction  between  pacemaker  and  heart,  shown  in  Fig.  4.  The ELT starts  with  an  early  ventricular   contraction  (PVC), which  is  a  common scenario even in a healthy  person. The  electrical  signal travels from the ventricle to  the  atrium  (red  arrows  in  Fig.  4),  triggering  an  atrial  sense (AS), i.e.  the pacemaker  senses an event in the right atrium.  As a result, the pacemaker  paces the ventricle (VP) after a pre-programed  delay (AVI), which triggers ventricle to

atrium conduction again and this VP  → AS  → VP  loop persists. The ventricular

rate  during  ELT is determined  by  the  conduction  delay  from the  ventricle  to the atrium  and the programed  delay  in the pacemaker,  which  is very  fast. The

healthy  condition  in Fig.  Fig.  4 demonstrates  the same input-output  sequence

ELT

A channel

V channel

PVC

AS	AS	AS	AS

Events

VS

Healthy

A channel

V channel

AVI

AVI	AVI

VP	VP	VP

9

Fig. 4. Endless-loop Tachycardia (ELT) and a healthy heart condition mapped to the

same input-output  execution of the pacemaker (middle sequence). The heart model

should have the details to resolve this ambiguity.

as ELT. In it, the pacemaker paces the ventricle after each atrial event (AS) generated by the SA  node (as opposed to an AS  conducted  from the ventricle  as in ELT). This  is correct  pacing  at an appropriate  rate  and maintains  adequate blood flow. The  heart model should be able to generate these two conditions to assess the pacemaker’s  ELT elimination capabilities  and that they do not affect other healthy  heart functions.

5   Multiscale Modeling

From  the example  above,  we can  see that  in order to perform model  checking on the  closed-loop  system,  the  heart  model  should  not only  cover  all  possible inputs  to  the  pacemaker   specified  in  the  requirement,  but  also  have  enough details  to  resolve  ambiguities  of executions  that  may  introduce  false-positives and false-negatives.  Fig.  5 shows a collection of heart conditions modeled by the EP  heart models. These models are definitely not exhaustive.  Model checking the pacemaker  model with all these heart models will not guarantee  absolute  safety. By  using physiological  abstraction  rules (R1-R7 in Fig.  5), the heart conditions can be generalized  and expand  the possible inputs to the pacemaker.  The  result

is a  heart  model  Hall  with  only  two  node automata  correspond  to the  inputs

to the pacemaker.  By  allowing  both  node automata  to be able  to send inputs

to the pacemaker  [0, ∞]msec after  the last  input,  the heart  model  Hall covers

all possible inputs to the pacemaker.  However,  Hall cannot distinguish the ELT

condition from the healthy condition due to the lack of representation of ventricle

to atrium  conduction.  The  heart  model Hcond   models the conduction  between

the  atrium  and  the  ventricle  with  a  path  automata,   thus  is  the  appropriate heart  model  to evaluate  ELT.  Similarly, more  complex  properties  will  lead  to appropriately detailed  models along the Model Abstraction Tree.

Heart

Conditions

Normal Sinus Rhythm

Bradycardia

1---------1

AV Block

Bundle Branch Block

Sinus

Tachycardia

Atrial

Flutter

AVNRT Atrial

Fibrillation

Premature

Ventricle

Contraction

$

R3

!'!o.$-"\

$"\:\	•

$

$--R2._..

V 	4V

-V-en-tr-icle---i_ - Tachycardia

1-

ibrillation	· ...!j,V

L-Ven-tri-cle    . _

Fig. 5. Multi-scale modeling of the heart. The heart model at a higher level  (further to the  right) contains all  possible inputs to the  pacemaker from  the  heart  models at previous levels

Model Checking

Simulation-based Testing

Platform Testing

•  Non-deterministic UPAAL Model

•  Abstract Heart and Device models

•  Deterministic Simulink Model

•  Increased complexity and interactions

• Hardware implementation of verified models

• Heart-on-a-Chip Platform

Fig. 6. Model translation framework. The pacemaker design is verified using model

checking and automatically  translated into code implementation.  Heart models are

available at all levels

6   From Verified Model  to Verified Code

During  model checking,  the abstract  model of the pacemaker  is verified against safety requirements. The abstract  pacemaker model is then automatically syn- thesized  into  simulation  models  and  into  a  code  implementation  (Fig.   6)  us- ing the UPP2SF tool we developed.  This  automatic  synthesis provides rigorous traceability throughout  the development  process  and  ensures that  the verified model  is translated  into verified  code.  Similarly   the  heart  model  is translated from timed  automata  to Simulink  and also synthesized  into a Heart-on-a-Chip for platform  level testing.

7   Model-Based Clinical Trials

The final step before the introduction of a medical device to market is usually the

clinical trial. The  randomized clinical trial (RCT) is considered to be the “gold

standard” for guaranteeing  that  a  medical  intervention  is safe and  efficacious

[4]. In the case  of new medical  devices  classified  as “significant  risk”,  an RCT is even mandated  by  the FDA before allowing  the device  on the market.  There are many  variations  of RCTs, but they  all share a basic  outline, an example  of which we now illustrate.

Suppose that a manufacturer  of medical devices is designing a new pacemaker

that’s  supposed  to assist  in treating  certain  abnormal  cardiac  rhythms,  or ar-

rhythmias. Both  the hardware  and software are tested by the company to ensure

it satisfies certain specifications, perhaps using model-based methods outlined above.  The  device may then be implanted  and tested on animals. But  up to this point,  the effect of the device  on humans has not been observed.  Observations of interest are not merely restricted  to whether the device  operates as intended or not.  We  want  to know  (and  regulators  require  evidence  of )  whether  it  can be implanted  safely; whether it has unexpected  side effects, or whose magnitude is too large;  and,  very  importantly,  whether  it treats  the targeted  arrhythmias better than current  medical  care,  thus justifying  its release on the market.

The  RCT seeks to answer  these questions by  comparing  two  groups  of pa-

tients: one which is implanted  with the investigational  device,  aka the treatment

group, and  one which  is on standard  medical  care,  aka  the control group. The

assignment  of a  patient  to the treatment  or control  group  is done at random,

which guarantees  the validity  of the statistical  tests used to analyze  the results,

and helps ensure that the two groups are comparable in terms of the physiolog-

ical  factors  that  might  affect the outcome  of the intervention.  Both  treatment and  control  groups  are  then  monitored  for a  pre-determined  amount  of time,

at the end of which  the rate  of treated  arrhythmias  is evaluated  in each  group (this  is the so-called  clinical  outcome  of the trial).  Finally, statistical  methods like the t-test  are applied  to determine whether the difference in rates between

the groups, if any,  is significant, i.e. is unlikely  to be due to chance  alone.

The  recognized  superiority  of RCTs comes at  a cost:  an RCT is in general a ma jor effort involving  patients,  medical  investigators,  ethics boards,  biostatis-

ticians, regulators,  companies, clinical centers, and large sums of money easily running in the millions of dollars. In addition, technical errors can arise at almost every  step of the trial planning,  jeopardizing  the validity  of the results. Finally, even if the trial is well-planned,  poor execution,  unexpected  events or even just pure chance  can lead to the wrong conclusions.

The  application  of computer  models to the medical  domain presented above have largely centered on the design, verification and deployment of a given device instance,  and have  mostly  eschewed  matters  related  to the clinical  trial.  There is however now an opportunity to use these computer models to assist in the planning and conduct of RCTs (see, e.g., the Avicenna  pro ject).  We call this the Model-Based  Clinical  Trial (MBCT).  Broadly speaking,  we define an MBCT  to

be a trial in which the sub jects are computer models of the biological phenomena

Fig. 7. Model-based Clinical Trials

being studied, including the effect of the device, rather than humans. An  MBCT

is not a replacement  for a clinical  trial:  rather, it will allow us to run very large-

scale  targeted  simulated  trials  to better  inform our conduct  of an actual  RCT. For example, we can study how variations in a patient’s physiological  parameters, like speed of propagation of the electric current in the heart, affects the safety and efficacy of the device. This is doable in MBCT and provides valuable  insight into which patients should be enrolled in the trial (and in whom the device is most efficacious). Another  application  would be to get tighter  estimates of statistical quantities  like  effect size  needed  before  the  conduct  of the  trial.  Unlike  drug trials, the results of medical device trials depend on the skill of the physician operating  the  device  or  implanting  it.  An  MBCT that  models  the  effects of physician  errors,  like  added  noise on dislodged  pacemaker  leads,  could  inform the trial investigators as to how much training is necessary for the physicians involved  in the trial.

For  an MBCT, it is not sufficient to only validate  that  the model structure

can produce physiologically correct behavior.  The model cohort as a whole must

also present the right variability to effectively be treated  as a group of patients. (The  model  cohort  is  the  group  of models  enrolled  in the  trial).  Our  current approach  is to learn the constraints on parameters  for the parameterized  model

using a sample of real patients’  data, as indicated  by markers   1   and   2   in Fig.

7. Ideally  this sample of patients is a cohort from a previous trial. These learned

constraints are then used to generate more instances of the model (marker    3 ), and  this  constitutes  our  model  cohort.  The  device  is then  connected  to  these models (marker   4 ) and the outcomes of interest are evaluated,  such as incidence of adverse  events.

Early  efforts tying  physiological  modeling  to clinical  trials  also  include  the

UVA/PADOVA diabetes  model  which  is used  to  generate  simulated  patients.

New diabetes  control algorithms  are evaluated  on simulated  patients  instead of animals  to assess their  efficacy.  These  applications  of MBCT usher a  new era of exciting  research  challenges  for the  closed-loop  verification,  validation  and testing medical  devices  at scale.  MBCT have the potential  to reduce the scope, cost and probability  of failure of clinical  trials  of medical  devices  with complex hardware  and  software.  MBCT when viewed  as a  rapid  certification  toolchain to speed up medical  device  approvals  is gaining increasing  traction  both within the regulatory  environment and the medical  device  industry.

References

[1] E.  Clarke, O. Grumberg, and D. Peled.  Model Checking. MIT  Press,  1999.

[2] Z.  Jiang,  M.  Pajic, and  R.  Mangharam.    Cyber-Physical Modeling  of Im-

plantable   Cardiac  Medical  Devices.    Proceedings of the IEEE, 100(1):122

–137, Jan. 2012.

[3] P.  L.  Jones.  Senior Systems/Software Engineer,  Office of Science  and Engi-

neering Laboratories,  US  FDA. Personal communication, 2010.

[4] L.  M.  Friedman   and  C.  D.  Furberg   and  D.  L.  DeMets.    Fundamentals of

Clinical Trials.  Springer,  2010.

[5] A.  Pullan,  L.  Cheng,  and M. Buist.  Mathematically Modelling the Electrical

Activity of the Heart: From Cell to Body Surface and Back Again.   World

Scientific,  2005.

[6] N.  A.  Trayanova and  P.  M.  Boyle.    Advances in  modeling  ventricular   ar-

rhythmias:  from mechanisms to the clinic.  Wiley Interdisciplinary Reviews:

Systems Biology and Medicine, 6(2):209–224, 2014.